STERILIZATION & DISINFECTION

Additional High-Yield Notes — Gap-Filler Edition

Supplement to your existing notes — covering the topics mock tests love to ask

Sources: Hugo & Russell | Ananthanarayan & Paniker | Lachman | Aulton | Indian Pharmacopoeia

For: GPAT | Drug Inspector | Pharmacist Recruitment | NIPER | University Exams

HOW TO USE THIS SUPPLEMENT This document deliberately focuses on the DETAILED, NUMERICAL and COMPARISON-BASED areas that are under-covered in most basic notes but are heavily tested in mock papers: • Sterilization kinetics with worked numericals (D, Z, F, F₀, SAL) • Every heat method (pasteurization, tyndallization, inspissation, autoclave types) • Radiation, gaseous (EO/formaldehyde/BPL/plasma) and filtration in exam-depth • Full disinfectant classification + mechanisms + concentrations • Evaluation of disinfectants (Rideal-Walker, Chick-Martin, Kelsey-Sykes) • Sterility testing (IP), biological & chemical indicators • 60+ previous-year style MCQs with explanations + quick-revision tables

# 1. KEY DEFINITIONS — THE FINE DISTINCTIONS

Examiners love the subtle differences between these terms. Learn them as pairs.

| Term | Precise Meaning | Key Exam Point |
| --- | --- | --- |
| Sterilization | Complete destruction / removal of ALL microorganisms including bacterial spores | An ABSOLUTE term — an object is either sterile or not (no 'partially sterile') |
| Disinfection | Destruction of vegetative pathogens (not necessarily spores) on inanimate objects | Directed at NON-LIVING surfaces; may not kill spores |
| Antisepsis | Destruction/inhibition of microbes on LIVING tissue (skin, mucosa) | Antiseptic = used on living tissue; Disinfectant = on inanimate objects |
| Sanitization | Reducing microbial load to a safe public-health level | Used for utensils, catering equipment (e.g., 99.9% reduction) |
| Decontamination | Removal/neutralization of contaminants so item is safe to handle | First step before disinfection/sterilization |
| Asepsis | Techniques/conditions that PREVENT entry of microbes | Medical asepsis vs Surgical asepsis (sterile field) |
| Sepsis / Sepsis | Presence of pathogens or their toxins in blood/tissue | Opposite of asepsis |
| -cidal vs -static | -cidal = KILLS microbes; -static = INHIBITS growth (reversible) | Bactericidal, sporicidal, virucidal, fungicidal vs bacteriostatic |
| Sterilant | A chemical agent capable of killing spores (achieving sterility) | e.g., glutaraldehyde, EO, H₂O₂ plasma, peracetic acid |
| Preservative | Substance preventing microbial growth in a formulation during shelf-life/use | Bacteriostatic in nature; e.g., benzalkonium Cl, parabens, phenylmercuric nitrate |

HIGH YIELD: The Resistance Ladder (most → least resistant) Prions > Bacterial spores > Coccidia (Cryptosporidium) > Mycobacteria > Non-enveloped (small) viruses > Fungi > Vegetative bacteria > Enveloped viruses → A method that kills bacterial SPORES will kill everything below them. → Prions resist normal autoclaving (need 134°C/18 min pre-vacuum + 1 N NaOH).

# 2. STERILIZATION KINETICS & THERMAL CONCEPTS

Microbial death by heat follows FIRST-ORDER (logarithmic) kinetics — a constant FRACTION (not number) of the population dies per unit time. Hence sterility is expressed as a PROBABILITY, never as an absolute zero.

## 2.1 The Core Parameters

| Parameter | Definition | Formula / Typical Value | Exam Note |
| --- | --- | --- | --- |
| D-value (Decimal Reduction Time) | Time (at a constant temp) to kill 90% of the population, i.e., reduce it by 1 log (factor of 10) | D = t / (log N₀ − log N) Unit: minutes | Higher D-value = MORE resistant organism. D depends on organism, temp & medium |
| Z-value | The temperature rise (°C) needed to reduce the D-value to 1/10th (one log) | Typically Z = 10°C for moist heat (steam) | Z reflects how sensitive D is to temperature; smaller Z = more temp-sensitive |
| F-value | Equivalent time (min) of heating at a stated temperature delivered to a product | F = D(base) × (log N₀ − log N) | Lethality expressed at a reference temperature |
| F₀-value | Equivalent lethality in minutes delivered at 121°C with Z = 10°C | Overkill standard: F₀ ≥ 12 minutes | The MOST asked value — basis of moist-heat cycle design |
| SAL (Sterility Assurance Level) | Probability of a single unit being NON-sterile after the process | Pharma standard SAL = 10⁻⁶ | '10⁻⁶' = not more than 1 non-sterile item per 1,000,000 items |
| Inactivation Factor (IF) | Factor by which population is reduced | IF = N₀ / N = 10^(t/D) | Links exposure time and D-value to total kill |

## 2.2 Thermal Death Terms (often confused)

| Term | Definition |
| --- | --- |
| Thermal Death Time (TDT) | Minimum TIME required to kill a microbial suspension at a GIVEN temperature |
| Thermal Death Point (TDP) | Lowest TEMPERATURE that kills all organisms in a suspension in 10 minutes |
| Decimal Reduction Time | Same as D-value (time for 1 log/90% reduction) |

WORKED NUMERICAL 1 — D-value Q: A population of 10⁶ spores is reduced to 10² after 8 minutes at 121°C. Find D₁₂₁. D = t / (log N₀ − log N) = 8 / (6 − 2) = 8 / 4 = 2 minutes. Interpretation: every 2 minutes at 121°C reduces the population by one log (90%).

WORKED NUMERICAL 2 — F₀ & sterility Q: Initial bioburden = 10² spores/unit, D₁₂₁ = 1.5 min. Time at 121°C to reach SAL 10⁻⁶? Total log reduction needed = log(10²) − log(10⁻⁶) = 2 − (−6) = 8 logs. Time = D × log reduction = 1.5 × 8 = 12 minutes (≈ the F₀ ≥ 12 overkill rule). This is exactly WHY the pharmaceutical overkill standard is F₀ = 12 minutes.

WORKED NUMERICAL 3 — Z-value Q: D₁₂₁ = 2 min and D₁₁₁ = 20 min. Find Z. D increased 10× (2 → 20) when temp fell by (121 − 111) = 10°C. A 10× change in D over 10°C ⇒ Z = 10°C.

# 3. MOIST HEAT STERILIZATION — ALL METHODS

Mechanism of kill: moist heat causes IRREVERSIBLE COAGULATION & DENATURATION of structural proteins and enzymes. Moist heat is far more efficient than dry heat (latent heat of steam + better penetration), so it kills at LOWER temperatures.

## 3.1 Temperature < 100°C

| Method | Conditions | Use / Notes |
| --- | --- | --- |
| Pasteurization – Holder (LTHT) | 63°C for 30 min | Milk; kills pathogens, NOT spores |
| Pasteurization – Flash (HTST) | 72°C for 15 sec | Rapid; better flavour retention |
| Pasteurization – UHT | 125–140°C for 1–2 sec | Long shelf-life 'sterile' milk |
| Inspissation | 80–85°C for 30 min × 3 days | Solidifies serum media (Loeffler, LJ medium) |
| Vaccine bath | 60°C for 1 hr | Heat-labile vaccine preparation |

HIGH YIELD: Why 63°C for pasteurization? Pasteurization temp is set to kill Coxiella burnetii (Q-fever) — the MOST heat-resistant non-sporing pathogen in milk. Also kills Mycobacterium bovis (TB), Brucella, Salmonella. Phosphatase test = checks adequacy of pasteurization (enzyme destroyed if done right).

## 3.2 Temperature at 100°C

| Method | Conditions | Notes |
| --- | --- | --- |
| Boiling | 100°C for 10–30 min | Kills vegetative forms; NOT reliably sporicidal |
| Tyndallization (Fractional / Intermittent) | 100°C (free steam) 20–30 min on 3 successive days | Between heatings, spores germinate → killed next day. For sugars/gelatin media that can't stand autoclaving |
| Steam at 100°C (Koch/Arnold) | Free flowing steam 90 min | Single exposure; not sporicidal |

## 3.3 Temperature > 100°C — The Autoclave

The autoclave works on the principle that water boils when its vapour pressure equals the surrounding pressure; increasing pressure raises the boiling point, giving SATURATED STEAM UNDER PRESSURE. It is the steam (not the pressure) that sterilizes.

| Temperature | Pressure (above atm.) | Holding Time |
| --- | --- | --- |
| 121°C | 15 psi (1.05 kg/cm²) | 15 minutes (standard) |
| 126°C | 20 psi (1.40 kg/cm²) | 10 minutes |
| 134°C | 30 psi (2.10 kg/cm²) | 3 minutes |
| 115°C | 10 psi (0.70 kg/cm²) | 30 minutes |

Holding time = penetration time + destruction time + safety margin (NOT just the clock time)

Air must be removed — trapped air lowers the true temperature of the steam (air is a poor conductor)

Downward-displacement (gravity) autoclave: steam pushes denser air out the bottom drain

Pre-vacuum (high-vacuum) autoclave: pump removes air first → faster steam penetration (porous loads, drapes)

Used for: aqueous solutions, dressings, surgical instruments, culture media, rubber

NOT used for: oils, powders, greases (steam can't penetrate) → use dry heat instead

AUTOCLAVE QUALITY CONTROL (very high yield) Biological indicator: Geobacillus stearothermophilus spores (kills at 121°C → proves sterility) Bowie-Dick test: checks AIR REMOVAL & steam penetration in pre-vacuum autoclaves (not lethality) Chemical indicators: Browne's tube (green = sterile), autoclave tape (stripes appear) Thermocouples: physical measurement of chamber/load temperature

# 4. DRY HEAT STERILIZATION

Mechanism of kill: dry heat destroys by OXIDATION of cell constituents, protein denaturation and toxic effect of raised electrolyte levels. It needs HIGHER temperatures and LONGER times than moist heat.

| Method | Conditions | Application |
| --- | --- | --- |
| Red heat | Heat to redness in flame | Inoculating loops, wires, forceps tips |
| Flaming | Pass through flame (no redness) | Scalpels, mouths of tubes, glass slides |
| Incineration | Complete burning to ash | Contaminated waste, carcasses, dressings |
| Hot Air Oven | 160°C / 2 h, 170°C / 1 h, 180°C / 30 min | Glassware, oils, powders, greases, metal instruments |
| Infrared radiation | Radiant heat (conveyor) | Rapid mass sterilization of syringes, catheters |

HOT AIR OVEN — EXAM POINTS Holding temperature/time (IP): 160°C for 2 hours is the classic combination Biological indicator: Bacillus atrophaeus (formerly B. subtilis var. niger) spores Must NOT overload; allow hot-air circulation; cool slowly to avoid glassware cracking Ideal for moisture-sensitive & steam-impenetrable items: fixed oils, glycerin, liquid paraffin, dusting powders, glassware (pipettes, flasks), sharp metal instruments

MOIST vs DRY HEAT — one-line comparison Moist heat kills by COAGULATION of proteins at LOWER temp (better penetration, latent heat). Dry heat kills by OXIDATION at HIGHER temp/longer time (for oils, powders, glassware).

# 5. RADIATION STERILIZATION

## 5.1 Ionizing Radiation ('Cold Sterilization')

| Feature | Details |
| --- | --- |
| Types | Gamma rays, X-rays, high-energy electron (cathode/beta) beams |
| Sources | Gamma: Cobalt-60 (chief), Caesium-137; Electron beam: linear accelerators |
| Sterilizing dose | 25 kGy (= 2.5 Mrad) is the accepted standard dose |
| Mechanism | Ionization → free radicals + direct DNA strand damage (high LET, deep penetration) |
| Advantages | No heat (ideal for heat-labile items), high penetration, done on final packed product |
| Applications | Disposable plastics (syringes, catheters), sutures, surgical blades, gloves, PPE, some drugs |
| Biological indicator | Bacillus pumilus spores (classical) |

RADIATION DOSE TERMS (food/pharma) Radappertization: high dose (25–50 kGy) → commercial 'sterility' Radicidation: lower dose to kill non-spore pathogens (like radiation pasteurization) Radurization: low dose to reduce spoilage organisms & extend shelf-life Order of resistance to radiation is NOT the same as to heat (e.g., some Deinococcus are radioresistant).

## 5.2 Non-Ionizing Radiation

| Type | Wavelength / Nature | Mechanism & Use |
| --- | --- | --- |
| Ultraviolet (UV) | 200–280 nm; germicidal peak ≈ 260–265 nm | Forms THYMINE DIMERS in DNA. Poor penetration → surface, air (BSC/laminar flow), water disinfection |
| Infrared (IR) | Longer than visible | Acts as DRY HEAT (thermal), not a true radiation kill |

UV — HIGH YIELD DNA absorption maximum ≈ 260 nm — this is why 254 nm mercury lamps are used. Damage: pyrimidine (thymine) dimers block replication. Repair: photoreactivation (light-dependent) & excision (dark) repair can reverse UV damage. Limitation: cannot penetrate glass, plastic, turbid liquids → surface/air only, not for deep sterilization.

# 6. GASEOUS (CHEMICAL VAPOUR) STERILIZATION

## 6.1 Ethylene Oxide (EO / EtO)

The most important low-temperature sterilant for heat- and moisture-sensitive devices. EO is an ALKYLATING agent — it alkylates sulphydryl (-SH), amino, carboxyl and hydroxyl groups of proteins and nucleic acids, blocking reproduction.

| Parameter | Requirement / Note |
| --- | --- |
| Nature | Colourless gas, sweet ethereal odour; FLAMMABLE & EXPLOSIVE in air |
| Made non-explosive by | Diluting with inert gas: CO₂ or nitrogen (e.g., 10% EO + 90% CO₂) |
| Concentration | ~450–1200 mg/L |
| Temperature | 30–60°C (typically 54°C) |
| Relative humidity | 30–60% (moisture essential — dry spores are resistant) |
| Exposure time | 2–48 h depending on load; then AERATION to remove residues |
| Toxic residues | Ethylene glycol & ethylene chlorohydrin → must aerate |
| Hazards | Toxic, mutagenic, carcinogenic, irritant → needs aeration & monitoring |
| Biological indicator | Bacillus atrophaeus spores |
| Uses | Plastics, catheters, sutures, prosthetic heart valves, ventilators, powders, complex devices |

## 6.2 Other Gaseous / Vapour Agents

| Agent | Key Facts |
| --- | --- |
| Formaldehyde (gas) | Alkylating; fumigation of rooms/safety cabinets (with KMnO₄ or steam). Formalin = 37–40% aqueous solution; paraformaldehyde on heating releases gas |
| Beta-propiolactone (BPL) | More active than formaldehyde but POOR penetration & CARCINOGENIC. Used to sterilize vaccines, sera, tissue grafts, blood plasma |
| H₂O₂ gas plasma (STERRAD) | Low-temp; free radicals from vaporized H₂O₂ in plasma phase. For heat/moisture-sensitive devices; no toxic residue |
| Peracetic acid vapour | Strong oxidizer, sporicidal even at low temp; used for endoscopes (STERIS) |
| Chlorine dioxide / Ozone | Oxidizing gases; ozone for water & some device sterilization |

GAS STERILANT MECHANISMS — remember this split ALKYLATING agents: Ethylene oxide, Formaldehyde, Beta-propiolactone. OXIDIZING agents: Hydrogen peroxide plasma, Peracetic acid, Ozone, Chlorine dioxide.

# 7. STERILIZATION BY FILTRATION

Filtration PHYSICALLY REMOVES microbes rather than killing them — the method of choice for HEAT-LABILE liquids (sera, antibiotic & enzyme solutions, vitamins, hormones) and for air. Note: ordinary filters do NOT remove viruses or mycoplasma.

| Filter Type | Material / Nature | Notes |
| --- | --- | --- |
| Membrane filter | Cellulose acetate/nitrate, PVDF, PTFE, nylon (screen filter) | Sterilizing grade = 0.22 µm; 0.45 µm for bioburden reduction. Reusable-checkable integrity |
| Sintered glass filter | Fused ground glass discs | Chemically inert, reusable, brittle |
| Seitz filter | Asbestos + cellulose pad (depth filter) | Adsorbs organisms; asbestos now largely obsolete |
| Candle filters | Diatomaceous earth (Berkefeld), unglazed porcelain (Chamberland), (Mandler) | Traditional depth filters for water/fluids |
| HEPA filter | Pleated glass-fibre | Removes 99.97% of particles ≥ 0.3 µm → laminar flow / BSC air |

FILTRATION — HIGH YIELD NUMBERS & TESTS Sterilizing membrane pore size = 0.22 µm (also written 0.2 µm). 0.45 µm does NOT guarantee sterility (some small organisms pass). HEPA = 99.97% efficiency at the most-penetrating size 0.3 µm. Integrity tests: Bubble point test, Pressure hold / diffusion test. Filter validation organism: Brevundimonas diminuta (Pseudomonas diminuta) for 0.22 µm; Serratia marcescens historically for 0.45 µm. Depth filter (traps within matrix) vs Screen/membrane filter (sieves at surface).

# 8. CHEMICAL DISINFECTANTS & ANTISEPTICS

Classify by chemical group — examiners ask 'which class does agent X belong to' and 'what is its mechanism'. Master the table below.

| Class | Examples | Mechanism | Uses / Notes |
| --- | --- | --- | --- |
| Phenolics | Phenol (carbolic acid – the standard), cresol (Lysol), chloroxylenol (Dettol), hexachlorophene | Membrane damage + protein denaturation | Surfaces; phenol is the reference for phenol coefficient |
| Alcohols | Ethanol (60–90%, optimum 70%), isopropanol | Protein denaturation + lipid dissolution (needs water) | Skin antiseptic; NOT sporicidal; volatile |
| Halogens – Chlorine | Hypochlorites (bleach), chloramine, Cl₂ gas | Oxidation (hypochlorous acid) | Water, surfaces; inactivated by organic matter |
| Halogens – Iodine | Tincture of iodine, iodophors (povidone-iodine/Betadine) | Oxidation / iodination of proteins | Skin & wound antiseptic; iodophors are less irritant |
| Aldehydes | Glutaraldehyde (2%, Cidex), formaldehyde, OPA (0.55%) | ALKYLATION of proteins/nucleic acids | Glutaraldehyde: cold sterilant/HLD for endoscopes (sporicidal) |
| Oxidizing agents | Hydrogen peroxide, peracetic acid, potassium permanganate | Free radical oxidation of cell components | H₂O₂ & peracetic acid are sporicidal |
| Heavy metals | Silver nitrate, silver sulfadiazine, mercurials (thiomersal), copper | Protein precipitation; OLIGODYNAMIC action | AgNO₃ (Credé's – ophthalmia neonatorum); burns (Ag-sulfadiazine) |
| Surface-active agents | Cationic: benzalkonium chloride, cetrimide, cetylpyridinium | Disrupt cell membrane (surfactant) | Cationic (QACs) most germicidal; anionic = soaps |
| Biguanides | Chlorhexidine (Savlon = chlorhexidine + cetrimide) | Membrane disruption + cytoplasm precipitation | Skin prep, mouthwash, surgical scrub |
| Dyes | Aniline dyes (brilliant/malachite green), acridine dyes (acriflavine, proflavine) | Interfere with nucleic acid / cell wall | Selective for Gram-positive; wound antiseptics |
| Acids / Alkalis | Benzoic, boric, sorbic acids; strong alkalis | pH-mediated protein/enzyme damage | Preservatives; food; laboratory decontamination |
| Gaseous agents | EO, formaldehyde, BPL (see Section 6) | Alkylation / oxidation | Heat-labile devices |

QUATERNARY AMMONIUM COMPOUNDS (QACs) — favourite MCQ trap Cationic surfactants (benzalkonium chloride, cetrimide) — most active surface-active class. INACTIVATED by: soaps/anionic detergents, hard water, organic matter, cotton/gauze. Poor against: Pseudomonas, Mycobacterium, bacterial spores, non-enveloped viruses. Pseudomonas can actually GROW in weak QAC solutions — a classic exam point.

AGENTS THAT ARE SPORICIDAL (i.e., true sterilants) Glutaraldehyde (2%), Formaldehyde, Ethylene oxide, Hydrogen peroxide, Peracetic acid, Chlorine dioxide, high-concentration hypochlorite. (Alcohols & QACs are NOT sporicidal.)

## 8.1 Factors Affecting Disinfectant Action

Concentration & contact time: higher concentration/longer time = greater kill (except alcohol needs water)

Temperature: activity generally increases with temperature

pH: affects the degree of ionization and hence activity of the agent

Organic matter (blood, pus, faeces): reduces activity (esp. chlorine, QACs) → clean before disinfecting

Microbial load (bioburden) & type: more organisms / spores = harder to kill

Nature of surface & presence of biofilm: biofilms are highly protective

# 9. EVALUATION OF DISINFECTANTS

This section is the SINGLE most common 'gap' in students' notes and a favourite for mock-test surprises. Learn the test names, organisms and what each measures.

| Test | Principle / What it Measures | Test Organism | Key Point |
| --- | --- | --- | --- |
| Phenol Coefficient – Rideal-Walker (RW) method | Ratio of the highest dilution of test disinfectant to that of phenol killing the organism in a set time (7.5 min but not 5 min) | Salmonella typhi | NO organic matter used → gives an over-optimistic (high) value |
| Chick-Martin method | Same idea BUT performed in presence of ORGANIC matter (yeast suspension / dried faeces) | Salmonella typhi (or S. aureus) | More realistic (lower) coefficient than RW |
| Kelsey-Sykes (Capacity) test | Capacity of a disinfectant to retain activity when repeatedly challenged with fresh culture | Ps. aeruginosa, Proteus, E. coli, S. aureus | In-use style; gives a recommended use-dilution |
| Use-dilution test (AOAC) | Determines maximum dilution that still kills on carriers (US official method) | S. choleraesuis, S. aureus, Ps. aeruginosa | Carrier (steel ring) method |
| In-use test | Tests the actual in-use diluted disinfectant from the ward for viable organisms | Environmental isolates | Detects contaminated / exhausted disinfectant in practice |

PHENOL COEFFICIENT — how to interpret Phenol coefficient = (dilution of disinfectant killing in time t) / (dilution of phenol killing in time t). PC > 1 → disinfectant is MORE potent than phenol. PC < 1 → LESS potent than phenol. PC = 1 → equal to phenol. Rideal-Walker uses S. typhi and NO organic load; Chick-Martin ADDS organic matter.

WORKED NUMERICAL — Phenol coefficient Disinfectant kills at a dilution of 1:300; phenol kills at 1:100 (same time & organism). PC = 300 / 100 = 3.0 → the disinfectant is 3× as potent as phenol.

# 10. STERILITY TESTING & VALIDATION (IP/USP)

Sterility testing confirms the ABSENCE of viable organisms in a sterilized product. Two official methods: (1) Membrane filtration (preferred) and (2) Direct inoculation.

| Medium | Detects | Incubation Temperature | Duration |
| --- | --- | --- | --- |
| Fluid Thioglycollate Medium (FTM) | Anaerobic + aerobic bacteria | 30–35°C | 14 days |
| Soybean-Casein Digest Medium (SCDM / TSB) | Fungi + aerobic bacteria | 20–25°C | 14 days |

Membrane filtration: sample filtered (0.45 µm), membrane washed & transferred to media → best for antibiotics/oils (removes inhibitors)

Direct inoculation: product added directly to media (used when filtration not feasible)

Incubation period: 14 days; observe for turbidity (growth)

Positive & negative controls must be run; aseptic technique / laminar flow essential

A single non-sterile unit → the batch fails (subject to retest rules)

## 10.1 Biological Indicators (memorize the organism per method)

| Sterilization Method | Biological Indicator (spore) |
| --- | --- |
| Moist heat / Autoclave (121°C) | Geobacillus stearothermophilus (formerly Bacillus stearothermophilus) |
| Dry heat / Hot air oven | Bacillus atrophaeus (formerly B. subtilis var. niger) |
| Ethylene oxide gas | Bacillus atrophaeus |
| Ionizing radiation | Bacillus pumilus |
| Filtration (validation) | Brevundimonas diminuta (Pseudomonas diminuta) |
| Low-temp H₂O₂ plasma | Geobacillus stearothermophilus |

## 10.2 Chemical Indicators

Bowie-Dick test: air removal / steam penetration test for pre-vacuum autoclaves

Browne's tube: liquid changes colour (red → green) when correct time-temperature reached

Autoclave/indicator tape: diagonal stripes develop — shows a pack was PROCESSED (not that it is sterile)

Integrating indicators (Class 5) & emulating indicators (Class 6): respond to all critical variables

CHEMICAL vs BIOLOGICAL INDICATOR — don't confuse Chemical indicator = shows the item was EXPOSED to the process (a colour/physical change). Biological indicator = uses resistant SPORES to PROVE microbial kill (the gold standard). 'Tape turned black' means processed, NOT necessarily sterile — classic MCQ trick.

# 11. PREVIOUS-YEAR STYLE MCQs (GPAT/Drug Inspector)

60+ MCQs focused on the gap topics above. Each has the answer and a short explanation.

Q1. The D-value of an organism is the time required to:

(a) Kill 100% (b) Reduce population by 90% (c) Reduce by 50% (d) Double the population

Answer: (b) Reduce population by 90% (1 log)

D-value (decimal reduction time) = time for a 1-log (10-fold) reduction at a constant temperature.

Q2. F₀ value is the equivalent lethality (min) delivered at:

(a) 100°C (b) 115°C (c) 121°C with Z=10°C (d) 134°C

Answer: (c) 121°C with Z = 10°C

F₀ is the reference lethality at 121°C assuming Z = 10°C; overkill standard is F₀ ≥ 12 min.

Q3. Sterility Assurance Level (SAL) for pharmaceutical products is:

(a) 10⁻³ (b) 10⁻⁶ (c) 10⁻¹ (d) 100%

Answer: (b) 10⁻⁶

SAL 10⁻⁶ = probability of not more than one non-sterile unit per million units.

Q4. Z-value is defined as the temperature change that alters D-value by:

(a) 2 times (b) 10 times (c) 50% (d) 100 times

Answer: (b) 10 times (one log)

Z-value = °C rise needed to reduce the D-value to one-tenth. For steam it is typically 10°C.

Q5. The standard sterilizing dose for radiation sterilization is:

(a) 2.5 kGy (b) 25 kGy (c) 250 kGy (d) 0.25 kGy

Answer: (b) 25 kGy

25 kGy (2.5 Mrad) is the accepted sterilizing dose for ionizing radiation.

Q6. The chief source of gamma radiation for sterilization is:

(a) Cobalt-60 (b) Uranium-235 (c) Radium-226 (d) Iodine-131

Answer: (a) Cobalt-60

Cobalt-60 is the principal gamma source; Caesium-137 is an alternative.

Q7. UV radiation kills microbes primarily by forming:

(a) Free chlorine (b) Thymine dimers (c) Ozone (d) Peroxide

Answer: (b) Thymine dimers in DNA

UV (peak ~260 nm) creates pyrimidine (thymine) dimers, blocking DNA replication.

Q8. Germicidal UV lamps emit maximally near:

(a) 190 nm (b) 254 nm (c) 365 nm (d) 400 nm

Answer: (b) 254 nm

Low-pressure mercury lamps emit ~254 nm, close to the DNA absorption max (~260 nm).

Q9. Ethylene oxide sterilizes by:

(a) Oxidation (b) Alkylation (c) Coagulation (d) Chelation

Answer: (b) Alkylation

EO alkylates -SH, -NH₂, -COOH and -OH groups of proteins & nucleic acids.

Q10. EO is made non-explosive by mixing with:

(a) Oxygen (b) CO₂ or nitrogen (c) Hydrogen (d) Methane

Answer: (b) CO₂ or nitrogen

Inert diluents (CO₂/N₂, formerly fluorocarbons) prevent EO's flammability/explosivity.

Q11. Toxic residue of ethylene oxide sterilization includes:

(a) Ethylene glycol & ethylene chlorohydrin (b) Formic acid (c) Methanol (d) Acetone

Answer: (a) Ethylene glycol & ethylene chlorohydrin

These residues necessitate a post-cycle aeration step.

Q12. Which humidity is required for effective EO sterilization?

(a) 0% (b) 30–60% (c) 90–100% (d) Humidity is irrelevant

Answer: (b) 30–60% RH

Moisture is essential; dry spores are highly resistant to EO.

Q13. Biological indicator for moist heat (autoclave) is:

(a) Bacillus atrophaeus (b) Geobacillus stearothermophilus (c) Bacillus pumilus (d) Clostridium tetani

Answer: (b) Geobacillus stearothermophilus

Its spores are highly heat-resistant and killed at 121°C, validating the autoclave.

Q14. Biological indicator for dry heat & EO is:

(a) G. stearothermophilus (b) Bacillus atrophaeus (c) Bacillus pumilus (d) E. coli

Answer: (b) Bacillus atrophaeus (formerly B. subtilis var. niger)

Same organism is used for dry heat and ethylene oxide validation.

Q15. Biological indicator for radiation sterilization is:

(a) B. pumilus (b) G. stearothermophilus (c) B. atrophaeus (d) S. aureus

Answer: (a) Bacillus pumilus

Classical biological indicator for ionizing radiation.

Q16. Autoclave standard conditions are:

(a) 100°C/10 psi/10 min (b) 121°C/15 psi/15 min (c) 160°C/2 h (d) 134°C/5 psi/3 min

Answer: (b) 121°C, 15 psi, 15 min

Saturated steam under pressure; it is the steam (latent heat), not pressure, that sterilizes.

Q17. What actually sterilizes in an autoclave?

(a) Pressure (b) Dry air (c) Saturated steam (d) Vacuum

Answer: (c) Saturated steam (moist heat)

Pressure only raises the boiling point; steam provides latent heat & penetration.

Q18. The Bowie-Dick test checks:

(a) Sterility (b) Air removal/steam penetration (c) Spore death (d) Pressure

Answer: (b) Air removal & steam penetration

It is a pre-vacuum autoclave function test, not a lethality/sterility test.

Q19. Hot air oven standard condition (IP):

(a) 121°C/15 min (b) 160°C/2 h (c) 100°C/30 min (d) 63°C/30 min

Answer: (b) 160°C for 2 hours

Also 170°C/1 h or 180°C/30 min; kills by oxidation.

Q20. Dry heat is preferred for:

(a) Aqueous injections (b) Oils, powders & glassware (c) Rubber (d) Culture media

Answer: (b) Oils, powders & glassware

Steam cannot penetrate oils/powders → dry heat used.

Q21. Pasteurization (holder method) conditions:

(a) 63°C/30 min (b) 100°C/10 min (c) 121°C/15 min (d) 72°C/30 min

Answer: (a) 63°C for 30 min (LTHT)

Flash/HTST = 72°C/15 sec; UHT = 125–140°C/1–2 sec.

Q22. Pasteurization temperature is set to kill which resistant pathogen?

(a) E. coli (b) Coxiella burnetii (c) Staph. aureus (d) Lactobacillus

Answer: (b) Coxiella burnetii

The most heat-resistant non-sporing pathogen in milk.

Q23. Tyndallization is:

(a) 100°C for 20–30 min on 3 successive days (b) 63°C/30 min (c) 160°C/2 h (d) 121°C/15 min

Answer: (a) Intermittent 100°C × 3 days

Also called fractional/intermittent sterilization; spores germinate between heatings then are killed.

Q24. Inspissation is used to sterilize/solidify:

(a) Water (b) Serum-containing media (Loeffler, LJ) (c) Oils (d) Glassware

Answer: (b) Serum media

80–85°C for 30 min on 3 successive days.

Q25. Sterilizing-grade membrane filter pore size:

(a) 0.45 µm (b) 0.22 µm (c) 1.2 µm (d) 5 µm

Answer: (b) 0.22 µm

0.45 µm does not guarantee removal of the smallest bacteria.

Q26. Filter validation is done using:

(a) E. coli (b) Brevundimonas diminuta (c) S. aureus (d) B. subtilis

Answer: (b) Brevundimonas (Pseudomonas) diminuta

Small size makes it the challenge organism for 0.22 µm filters.

Q27. HEPA filters remove particles ≥ 0.3 µm with efficiency:

(a) 90% (b) 95% (c) 99.97% (d) 100%

Answer: (c) 99.97%

HEPA supplies particle-free air to laminar-flow benches and BSCs.

Q28. Membrane filter integrity is checked by:

(a) Bubble point test (b) Gram staining (c) pH test (d) Bowie-Dick test

Answer: (a) Bubble point test

Also pressure-hold/diffusion tests confirm integrity.

Q29. The standard reference disinfectant for phenol coefficient is:

(a) Cresol (b) Phenol (c) Ethanol (d) Chlorhexidine

Answer: (b) Phenol

Phenol coefficient compares a disinfectant's potency to phenol.

Q30. Rideal-Walker method uses which test organism?

(a) S. aureus (b) Salmonella typhi (c) Ps. aeruginosa (d) E. coli

Answer: (b) Salmonella typhi

RW test is performed WITHOUT added organic matter.

Q31. Which test evaluates disinfectants in presence of organic matter?

(a) Rideal-Walker (b) Chick-Martin (c) Bubble point (d) Bowie-Dick

Answer: (b) Chick-Martin

Organic matter (yeast/faeces) gives more realistic (lower) coefficients.

Q32. A phenol coefficient of 3 means the disinfectant is:

(a) 1/3 as active as phenol (b) 3 times as active as phenol (c) Equal to phenol (d) Inactive

Answer: (b) 3 times as active as phenol

PC > 1 = more potent than phenol.

Q33. The capacity test for disinfectants is the:

(a) Kelsey-Sykes test (b) Rideal-Walker test (c) AOAC test (d) Koch test

Answer: (a) Kelsey-Sykes (capacity) test

Assesses ability to retain activity under repeated microbial challenge.

Q34. Optimum bactericidal concentration of ethanol is:

(a) 100% (b) 70% (c) 40% (d) 10%

Answer: (b) 70%

Water is needed for protein denaturation; absolute alcohol is less effective.

Q35. Which agent is NOT sporicidal?

(a) Glutaraldehyde (b) Ethylene oxide (c) 70% ethanol (d) Peracetic acid

Answer: (c) 70% ethanol

Alcohols kill vegetative cells but not spores.

Q36. 2% glutaraldehyde (Cidex) is chiefly used to:

(a) Sterilize glassware (b) Cold-sterilize endoscopes (c) Purify water (d) Preserve vaccines

Answer: (b) Cold-sterilize endoscopes

High-level disinfectant / chemical sterilant; sporicidal on long contact.

Q37. Which disinfectant class do benzalkonium chloride and cetrimide belong to?

(a) Phenolics (b) Quaternary ammonium (cationic surfactants) (c) Aldehydes (d) Halogens

Answer: (b) Quaternary ammonium compounds

Cationic surfactants; membrane-disrupting.

Q38. QACs are INACTIVATED by:

(a) Hard water & soaps (b) Sunlight only (c) Sugar (d) Nothing

Answer: (a) Hard water, soaps/anionic detergents & organic matter

Also weak against Pseudomonas, mycobacteria & spores.

Q39. Chlorhexidine belongs to which class?

(a) Biguanide (b) Aldehyde (c) Phenol (d) Alcohol

Answer: (a) Biguanide

Savlon = chlorhexidine + cetrimide; used as skin/mucosal antiseptic.

Q40. Oligodynamic action is shown by:

(a) Heavy metals (silver, copper) (b) Alcohols (c) Phenols (d) Aldehydes

Answer: (a) Heavy metals

Trace amounts of metal ions (Ag, Cu) are antimicrobial — oligodynamic action.

Q41. Silver nitrate (Credé's method) prevents:

(a) Tetanus (b) Ophthalmia neonatorum (c) TB (d) Rabies

Answer: (b) Ophthalmia neonatorum

1% AgNO₃ instilled into newborn eyes against gonococcal conjunctivitis.

Q42. Povidone-iodine is an example of a(n):

(a) Iodophor (b) Phenol (c) Biguanide (d) Alcohol

Answer: (a) Iodophor

Iodine complexed with a carrier (povidone); less irritant, sustained release.

Q43. Beta-propiolactone (BPL) is mainly limited by:

(a) Being non-toxic (b) Carcinogenicity & poor penetration (c) High cost only (d) Low activity

Answer: (b) Carcinogenicity & poor penetration

Despite high activity; used for vaccines, sera, tissue grafts.

Q44. Formalin is:

(a) 37–40% aqueous formaldehyde (b) 2% glutaraldehyde (c) 70% ethanol (d) 10% phenol

Answer: (a) 37–40% aqueous formaldehyde

Used for fumigation and tissue fixation.

Q45. STERRAD system uses:

(a) Hydrogen peroxide gas plasma (b) Steam (c) Dry heat (d) UV

Answer: (a) H₂O₂ gas plasma

Low-temperature sterilization with no toxic residue.

Q46. Fluid thioglycollate medium in sterility testing detects:

(a) Only fungi (b) Aerobic + anaerobic bacteria (c) Viruses (d) Prions

Answer: (b) Aerobic + anaerobic bacteria

Incubated at 30–35°C for 14 days.

Q47. Soybean-casein digest medium is incubated at:

(a) 20–25°C (b) 30–35°C (c) 37°C (d) 55°C

Answer: (a) 20–25°C

For fungi and aerobic bacteria; 14-day incubation.

Q48. The preferred sterility-testing method for antibiotic solutions is:

(a) Direct inoculation (b) Membrane filtration (c) Pour plate (d) Streak plate

Answer: (b) Membrane filtration

It removes inhibitory antibiotic before culture.

Q49. Sterility test incubation period is:

(a) 24 h (b) 48 h (c) 7 days (d) 14 days

Answer: (d) 14 days

Standard incubation to detect slow-growing contaminants.

Q50. Which is the MOST resistant to sterilization?

(a) Enveloped virus (b) Vegetative bacteria (c) Prions (d) Fungi

Answer: (c) Prions

Prions resist routine autoclaving; need 134°C/18 min + NaOH.

Q51. Among microbes, bacterial endospores are killed by all EXCEPT:

(a) Autoclaving (b) Ethylene oxide (c) 70% alcohol (d) Glutaraldehyde

Answer: (c) 70% alcohol

Alcohol is not sporicidal.

Q52. Incineration is best suited for:

(a) Reusable instruments (b) Contaminated waste/carcasses (c) Milk (d) Injections

Answer: (b) Contaminated waste/carcasses

Complete destruction by burning to ash.

Q53. Red heat is used to sterilize:

(a) Bacteriological loops & wires (b) Scalpels (c) Glass pipettes (d) Rubber

Answer: (a) Loops & wires

Heated to redness directly in flame.

Q54. Which chemical indicator only shows a pack was 'processed'?

(a) Autoclave tape (b) Spore strip (c) Thermocouple (d) Bowie-Dick pack

Answer: (a) Autoclave tape

Colour/stripe change indicates exposure, NOT sterility.

Q55. First-order (logarithmic) death kinetics implies:

(a) Constant number die per unit time (b) Constant fraction die per unit time (c) All die at once (d) None die

Answer: (b) Constant fraction per unit time

Hence sterility is expressed as a probability (SAL).

Q56. Peracetic acid is classed as a(n):

(a) Alkylating agent (b) Oxidizing agent (c) Cationic surfactant (d) Dye

Answer: (b) Oxidizing agent

Sporicidal even at low temperature; used for endoscopes.

Q57. Which pair correctly matches agent → mechanism?

(a) Glutaraldehyde → oxidation (b) EO → alkylation (c) H₂O₂ → alkylation (d) Alcohol → chelation

Answer: (b) Ethylene oxide → alkylation

Glutaraldehyde/formaldehyde/EO alkylate; H₂O₂/peracetic acid oxidize.

Q58. The gas used with EO to prevent explosion was historically a fluorocarbon; now common is:

(a) O₂ (b) CO₂ or N₂ (c) H₂ (d) He

Answer: (b) CO₂ or N₂

Inert diluent gases render the mixture non-flammable.

Q59. Which method removes rather than kills microbes?

(a) Autoclaving (b) Filtration (c) Radiation (d) EO gas

Answer: (b) Filtration

Physical removal; ideal for heat-labile fluids.

Q60. Depth filters differ from membrane filters in that they:

(a) Sieve only at the surface (b) Trap organisms within the filter matrix (c) Are always 0.22 µm (d) Cannot be used for liquids

Answer: (b) Trap within the matrix

Sintered glass/Seitz/candle filters are depth filters; membranes are screen filters.

Q61. Which is a low-temperature STERILANT (not just a disinfectant)?

(a) 70% ethanol (b) Benzalkonium chloride (c) Ethylene oxide (d) Cetrimide

Answer: (c) Ethylene oxide

EO achieves sterility (sporicidal); alcohols/QACs do not.

Q62. Ozone and chlorine dioxide sterilize by:

(a) Alkylation (b) Oxidation (c) Coagulation (d) Filtration

Answer: (b) Oxidation

Both are oxidizing gaseous agents.

# 12. QUICK REVISION — LAST-MINUTE TABLES

## Numbers You MUST Remember

| Parameter | Value |
| --- | --- |
| Autoclave (standard) | 121°C, 15 psi, 15 min |
| Autoclave (rapid) | 134°C, 30 psi, 3 min |
| Hot air oven | 160°C / 2 h (or 170°C/1 h, 180°C/30 min) |
| Pasteurization (holder/LTHT) | 63°C / 30 min |
| Pasteurization (flash/HTST) | 72°C / 15 sec |
| Pasteurization (UHT) | 125–140°C / 1–2 sec |
| Tyndallization | 100°C, 20–30 min × 3 days |
| Inspissation | 80–85°C, 30 min × 3 days |
| Radiation dose | 25 kGy (2.5 Mrad) |
| UV germicidal peak | 254–265 nm |
| Sterilizing filter pore size | 0.22 µm |
| HEPA efficiency | 99.97% at 0.3 µm |
| F₀ (overkill) | ≥ 12 min |
| SAL (pharma) | 10⁻⁶ |
| Sterility test incubation | 14 days |

## Biological Indicators — rapid recall

| Method | Indicator Organism |
| --- | --- |
| Moist heat / autoclave | Geobacillus stearothermophilus |
| Dry heat & Ethylene oxide | Bacillus atrophaeus |
| Ionizing radiation | Bacillus pumilus |
| Filtration validation | Brevundimonas diminuta |

## Mechanism Buckets

GROUP BY MECHANISM COAGULATION of protein: moist heat (autoclave, pasteurization). OXIDATION: dry heat, H₂O₂, peracetic acid, ozone, chlorine, iodine. ALKYLATION: ethylene oxide, formaldehyde, glutaraldehyde, beta-propiolactone. MEMBRANE disruption: alcohols, phenols, QACs, chlorhexidine. DNA damage: ionizing radiation, UV (thymine dimers). PHYSICAL removal: filtration.

## Evaluation Tests — rapid recall

| Test | Organism / Feature |
| --- | --- |
| Rideal-Walker | S. typhi; NO organic matter (over-estimates) |
| Chick-Martin | S. typhi/S. aureus; WITH organic matter (realistic) |
| Kelsey-Sykes | Capacity test under repeated challenge |
| Use-dilution (AOAC) | Carrier method; S. choleraesuis, S. aureus, Ps. aeruginosa |
| In-use test | Checks the actual diluted disinfectant in wards |

EXAM STRATEGY — highest-yield for THIS topic 1. D / Z / F / F₀ / SAL definitions & a simple numerical — asked almost every paper. 2. Biological indicator organism for each method — pure recall marks. 3. Autoclave vs hot air oven conditions & what each is used for. 4. Phenol coefficient (RW vs Chick-Martin) — the classic 'surprise' MCQ. 5. EO parameters, alkylation, residues, non-explosive mixing. 6. Filter pore size (0.22 µm), HEPA 99.97%, validation organism. 7. Which agents are sporicidal vs not (alcohols/QACs are NOT). 8. Pasteurization temps + Coxiella burnetii; tyndallization vs inspissation. 9. Sterility testing media (FTM vs SCDM), temperatures, 14-day incubation. 10. Resistance ladder: prions > spores > mycobacteria > ... > enveloped viruses.

— End of Supplement —  Best of luck for your exam!